## List1
| Parameter | Interpretation | Example Values, DEC (HEX) | Notes
Page 4, Byte 0 | Pulse width HB | Pulse width in chunks of 50 us, high byte
Page 4, Byte 1 | Pulse width LB | Pulse width in chunks of 50 us, low byte | "1" = 50 us; "2" = 100 us; "3" = 150 us, etc. | Step 50 us
Page 4, Byte 2 | Pulse Frequency HB | Pulse frequency timer byte word, high byte | "93" (0x5D) | 0x5D3D = 20 Hz
Page 4, Byte 3 | Pulse Frequency LB | Pulse frequency timer byte word, low byte | "61" (0x3D) |  | Page 5 (Duty Cycle programming) |  |  | Current amplitude programming
Page 5, Byte 0 | Stimulation Time On (Duty cycle), msec byte word | Milliseconds setting, frequency timer reload first ("lowest") byte | "0" | 0x027000 ≈ 4.875 sec; 0x0F0000 = 30 sec | Reload byte word | Seconds | Data array values | Desired stimulation current (uA): | 128
Page 5, Byte 1 | Stimulation Time On (Duty cycle), sec byte word, HB | Seconds setting, real-time-clock timer reload Low byte | "112" (0x70) |  | 0x008000 | 1 | Byte 0: "0", Byte 1: "0", Byte 2: "128" | Page 6, Byte 0: | 4.0716
Page 5, Byte 2 | Stimulation Time On (Duty cycle), sec byte word, LB | Seconds setting, real-time-clock timer reload High byte | "2" (0x02) |  | 0x027000 | 5 | Byte 0: "0", Byte 1: "112", Byte 2: "2"
Page 5, Byte 3 | On/Off byte | Enable/Disable stimulation circuitry | 0 or 1 | 0 = Off; 1 = On | 0x050000 | 10 | Byte 0: "0", Byte 1: "0", Byte 2: "5"
[5 empty] | 0x0A0000 | 20 | Byte 0: "0", Byte 1: "0", Byte 2: "10"
Page 6, Byte 0 | Current Amplitude | Current reference setting, allows control of the amplitude | "63" (0x3F) = 2.012 mA, step 31.4375 uA | Iload = Iref * 4; Iref = [0; 63] = [0 uA; 503 uA] => Iload = [0; 63] = [0; 2.012 mA] | 0x0F0000 | 30 | Byte 0: "0", Byte 1: "0", Byte 2: "15"
Page 6, Byte 1 | Stimulation Mode | Switch between available stimulation modes | 1 = channel scanning, starting at the pair 14-22, ending with 1-19, cycling at frequency set by the Page 5 parameters; 2 = single channel mode |  | 0x190000 | 50 | Byte 0: "0", Byte 2: "0", Byte 3: "25"
Page 6, Byte 2 | Stimulation Channel (For Single-Channel Stim) | Stimulation channel select in the single-channel mode | "0" (0x00) to "14" (0x0E) | Use HEX for channel enumeration | 0x1E0000 | 60 | Byte 0: "0", Byte 1: "0", Byte 2: "30"
Page 6, Byte 3 |  |  |  |  | 0x3C0000 | 120 | Byte 0: "0", Byte 1: "0", Byte 2: "60"
Page 8, Byte 0 | Current Stimulation Channel (Channel Scanning Mode) | Feedback from the device: which channel is being stimulated now | "0" (0x00) to "14" (0x0E)
Page 8, Byte 1
Page 8, Byte 2
Page 8, Byte 3